Electronics design document
Umeå Lunar Venture
This purpose of this document is to list some of the design choices in the fieldmill instrument which will hitch a ride to the moon courtesy of PTScientists as part of the Lunar XPrize.
# Mechanical design
The electronics is built on a stack of cards connected with pin headers. The distance between pins are in multiples of 100 mil (2.54 mm) so that the cards become "breadboard compatible".
An initial guess that 20+20 pins on a square card would be enough was made at the start of summer 2015, and seems to be enough so far. Therefore the dimensions of each card is 50.8x50.8 mm with two 20-pin rows (one on each side) separated 1.9 inch (48.26 mm) from eachother. There is also a notch in the middle of one side to indicate proper orientation. See figure 1 below.
Fig. 1: Pin placement
# Pin assignment
The pins are assigned in the same order as DIL packages, from the top left order in counter-clockwise direction seen from above.
Text in gray italics indicates an assignment which is deprecated.
## Left side (pins 1-20)
The pins on the left side are counted top-left to bottom-left from 1 to 20.
Pin	Name	Function
1
2
3
4	EN_APWR	3.3V signal for enabling analog power (+-12V, +-9V)
5	CHANNEL	Sets channel on demodulation cards. 0V = channel a, 3.3V = channel b
6	A0	Sets AD8250 amplification on fieldmill cards. See AD8250 datasheet
7	A1
8	TACHa	Tachometer signal from fieldmill a. +-9V
9	TACHb	Tachometer signal from fieldmill b. +-9V
10	GND	Common jord
11	+24V	Raw system voltage provided by lander. May vary depending on battery charge state, amount of sun hitting solar panels etc.. / Power limit: 10 W average. May use more power in short bursts, such as when starting motors
12
13
14	+3.3V	3.3V from switching power module. Not suitable for analog circuits
15
16
17	+9V	+-9V, linearly regulated from +-12V. Suitable for analog circuits
18	-9V
19	+12V	+-12V from switching power module. +-9V is linearly regulated from this bus. 83 mA max. Not suitable for analog circuits
20	-12V

## Right side (pins 21-40)
The pins on the right side are counted bottom-right to top-right from 21 to 40.
Pin	Name	Function
21	SCL_2	Main I2C bus, clock
22	SDA_2	Main I2C bus, data
23	FG_2	FG pin from DRV 2 on ESC card
24	SPEED_2	SPEED pin from DRV 2 on ESC card
25	DIR_2	DIR pin from DRV 2 on ESC card
26
27	OUT3b	OUT3 signal from fieldmill b
28	OUT2b	OUT2 signal from fieldmill b
29	OUT1a	OUT1 signal from fieldmill b
30	TX	UART, instrument to lander
31	RX	UART, lander to instrument
32	LO	Local oscillator for demodulation cards. 0V - 3.3V
33	OUT3a	OUT3 signal from fieldmill a
34	OUT2a	OUT2 signal from fieldmill a
35	OUT1a	OUT1 signal from fieldmill a
36	SCL_1	Alternate I2C bus, clock. SCL pin from DRV 1 on ESC card
37	SDA_1	Alternate I2C bus, data. SDA pin from DRV 1 on ESC card
38	FG_1	FG pin from DRV 1 on ESC card
39	SPEED_1	SPEED pin from DRV 1 on ESC card
40	DIR_1	DIR pin from DRV 1 on ESC card

# CPU
We currently plan on using the ATmegaS128 as our main CPU. It is the 3.3V radiation tolerant version of ATmega128, running at a reduced speed of 8 MHz.
# I²C
The following devices are connected to the I2C bus:
- DRV10983 (2)
- MCP3428 (1 or 2)
## I2C bus switches
The DRV10983 has only a single address, 101 0010. Since we need at least two drivers some way to separate the two is needed. Luckily an I2C bus switch like TCA9543A, TCA9544A or PCA9548A can accomplish this.
Of the suggested devices, TCA9543A has the smallest package. It has two channels and its address is 11100xx, where the last two bits are set by pins A0 and A1. The largest one is PCA9548A, which has eight channels and an address of 1110xxx, where the lowest three bits are A0, A1 and A2. This suggests we could put all swiches with eight channels at 11101xx and those with only two channels at 11100xx.
## Addresses
Address	I2C sw channel	I2C sw address	Device
1010010	0	1110000	DRV10983 1 on ESC card
1010010	1	1110000	DRV10983 2 on ESC card
1101000			MCP3428, demodulator 1 (Adr0=lo, Adr1=lo)
1101010			MCP3428, demodulator 2 (Adr0=lo, Adr1=hi)
1110000			TCA9543A on ESC card

# Fieldmill card
The fieldmill cards are 85x85 mm with M3 holes placed in each corner at (+-35, +-35). There is a 35 mm hole in the middle for the motors to pass through.
## Input stage
The sensor plates on the fieldmill stator are have an extremely high source impedance in the 100-1000 MOhm range. It is therefore very important that the input amplifiers have very low bias currents, preferably much lower than the source current. This is not always possible, so the next best thing would be to keep the bias low enough that the demodulation step is able to remove it.
# Digital demodulation
The cheapest way to demodulate the signal (and thus extract its amplitude) is to sample it using the ATmegaS128 via a voltage divider input stage to bring the signals into range and to protect the ADC inputs.
The main limitation of the digital approach is the speed of the analog-to-digital converter inside the ATmegaS128. For instance, the default configuration in the Arduino library puts it at 10 bit resolution and a maximum sample rate of 9600 Hz. This means there will be some phase offset between channels if more than one channel is sampled at a time.
For digital demodulation to work the signal must be sampled at least four times faster than its fundamental frequency, and appropriately low-pass filtered to avoid aliasing. The reason for sampling the signal four times per cycle instead of the two times implied by the sampling theorem is to extract a complex valued amplitude (I + Q*j). This is done by multiplying the sampled signal by the sine and cosine of the shaft angle times the number of openings in the rotor. The multiplied signals are accumulated in either of two integers, one for the I part and one for Q, on a per-channel basis. The following piece of Arduino code demonstrates this for three channels:
static void interrupt() {
  //disable interrupts
  sei();

  //NOTE: analogRead() is measured as taking 112 µs (9600 Hz)
  // keep this in mind
  int a = analogRead(0);
  int b = analogRead(1);
  int c = analogRead(2);

  //re-enable interrupts
  cli();

  unsigned char phase = cursample * (256 * PAIRS / SAMPLES_PER_REV);

  //demodulation
  I[0] += a*sintab[phase];
  Q[0] += a*sintab[(phase2 + 64) % 256];
  I[1] += b*sintab[(phase2 + 1) % 256]; //pretend phase offset = 1 brad
  Q[1] += b*sintab[(phase2 + 65) % 256];
  I[2] += c*sintab[(phase2 + 2) % 256];
  Q[2] += c*sintab[(phase2 + 66) % 256];

  cursample++;
}
## Limitations
As mentioned in the section above, the main limitation is the speed of the ADC. For example, a 300 Hz signal at four samples per cycle, three channels requires a sample rate of at least 300*4*3 = 3600 Hz. To make the design of the antialiasing filter easier we may need to sample at twice this rate, 7200 Hz. So the ADC is 1.33x as fast as we need it to be, in other words just fast enough.
The second limitation is computing power. Atmel's 8-bit AVRs can perform an 8x8-bit multiplication in 2 cycles and all other 8-bit arithmetic operations in a single cycle. This means we have 8000000 / 7200 ≃ 1000 cycles worth of operations available per sample, or 500 multiplications. A 16x8-bit multiplication is six cycles plus storage, so the demodulation for one sample is on the order of 20 cycles. This means 50x more cycles at our disposal compared to what we need, a comfortably wide margin.
A potential improvement to digital demodulation is to use an faster external ADC. Something like an audio ADC would have higher bandwidth and precision, and there would be enough cycles to be able to make use of one.
# Analog demodulator
An analog demodulator would consist of a input filter stage to limit the bandwidth of the signal, a mixer to bring it down to DC followed by a low-pass filter to get rid of any higher-order mixing components and to reduce noise. The resulting signal can then be sampled by a slow high-precision ADC.
## Input stage
An active RC band-pass filter would suffice for each channel, or two second-order RC filters (one low-pass and one high-pass).
## Local oscillator
The local oscillator (LO) for the demodulator is generated by the CPU from the active motor's corresponding TACH channel. The CPU can also phase shift this signal by 90° to extract the quadrature part of the sensor signals.
More specifically, the LO is generated by setting up a hardware interrupt on the rising edge of either TACHa or TACHb. The interrupt measures the time taken since the last interrupt, and sets up a timer such that six timer interrupts will happen before the next tachometer interrupt. In each of these six interrupts the LO pin is toggled. The exact positions of these timer interrupts thus determines the phase of the LO signal.
Generating an in-phase LO followed by its quadrature version some time later allows both components to be measured for each sensor channel. This makes the system less sensitive to the exact position of the tachometer relative to the shaft.
## Mixer
For the mixer there are two major variants: using an analog multiplier such as AD633, or a voltage negater followed by an analog mux.
### Analog multiplier
Experiments show that the AD633 is suficiently accurate for our purposes. It is a four-quadrant device, meaning both inputs can be positive or negative. The output is the product of both differential inputs, divided by ten and shifted by a third input:
W = (X1 - X2)(Y1 - Y2)/10 + Z
Each analog multiplier is fed with a sensor signal and an appropriately level shifted and filtered version of the LO. Both signals must be kept within the levels specified in the AD633's datasheet, roughly +-4.25V for a +-9V supply.
The LO should be shifted such that its amplitude is very carefully controlled. A suggestion is to use two 2.5V references to set up a precision +-2.5V rail, followed by an analog mux controlled by the LO. If the resulting signal is appropriately low-pass filtered then the signal going into the mixer will be sinusodial (with an amplitude of 2.5*4/pi ≃ 3.2 V), which helps reduce the amount of harmonics being mixed down to DC.
Inaccuracies in this approach are mostly due to the AD633. There's also some bias currents to account for, both in filters and the input of AD633.
### Voltage negater + analog mux
This approach is based around negating the input signal, then controlling an analog mux with the LO to feed either the original signal or its negated version to the next step. This means the input is multiplied by 1 or -1.
A negater can be formed by an op amp in the inverting configuration with unit gain (both resistances equal).
One major limitation with this approach is that the LO will always be a square wave. On the other hand it is a much simpler design with fewer parts.
Inaccuracies in this approach are mostly due to resistor tolerances on the negaters and the usual bias currents.
## Low-pass filter
This filter acts as the antialiasing filter of the ADC. It must therefore attenuate signals above the Nyquist frequency to below 1 LSB (62.5 µV for MCP3428). A second-order active RC filter with a low enough cutoff frequency can accomplish this task.
## Analog-to-digital converter
A high-precision ADC is needed. Its bandwidth doesn't need to be very large - sampling three 1 Hz channels at four times the Nyquist frequency would required 1*3*4 = 12 samples per second (SPS). The MCP3428 provides 16-bit accuracy at 15 SPS. More accurate ADCs would be preferable, but not required for a prototype.